## Slide 1: Removal of Motion Artifacts from Photoplethysmogram Data
By: Supervisor:
Salaar Ali Assad - 21100087 Dr. Muhammad Tahir
Sheraz Hassan - 21100285
Mohammad Afaq - 21100214
Department of Electrical Engineering
Lahore University of Management Sciences

## Slide 2: Recap – SPROJ 1

## Slide 3: Motivation
19% of the Pakistani population suffers from Cardiovascular Diseases (CVDs), making it the leading non-communicable disease in the country [1]
Increasing demand for Electrocardiogram (ECG) which are uncomfortable, expensive, limited and require medical supervision [2]
Wrist worn Photoplethysmogram (PPG) devices offer a simpler, cheaper and more comfortable alternative

## Slide 4: High-End Device
Empatica E4
4 LED’s and 2 photodiodes
PPG sensor (64HZ)
3-axis accelerometer (32Hz)
Temperature sensor
Pre-processing algorithm
Front and Back View of the Empatica E4

## Slide 5: Low-End Device
Maxim MAXREFDES103#
2 LED’s and 2 photodiodes
IR emitter
Variable sampling rate
Front and Back View of the Maxim MAXREFDES103

## Slide 6
6
MAXREFDES103 Signal = Perfusion waveform + PPG waveform + noise (2)
E4 Signal = PPG waveform + noise (1)
Low-end vs High-end

## Slide 7: Perfusion
Variation of blood flow volume in capillary network
Causes a baseline shift in the raw PPG signal.
PPG Signal without perfusion (top) and with perfusion (bottom) [4]

## Slide 8: Perfusion – Removal

## Slide 9
9
Raw PPG Signal (Blue) along with the approximated Perfusion Waveform (Red)
PPG Waveform obtained along with its Frequency Plot

## Slide 10: Breath Rate
Respiratory signal dominated by frequencies at 0.3 Hz
Apply band pass filter
Find Peaks
Average of distance between successive peaks in samples

## Slide 11: SPROJ 2

## Slide 12: Block Diagram

## Slide 13: Data Collection

## Slide 14: Classification
Rest position
Walking
Slight Arm Movement – Dominant
Slight Arm Movement – Non-Dominant
Typing
Writing
* Data not taken under clinical standards

## Slide 16: Motion Artifact Removal

## Slide 17: Existing Methods
Independent Component Analysis (ICA)
Adaptive Filtering with Accelerometer Signal
Deep Learning

## Slide 18: Attempted Approaches
Adaptive Filtering with Accelerometer Signal
Time Delay Neural Network
Weiner Filtering

## Slide 19: Our Approach

## Slide 20: Detection
Skewness
Kurtosis

## Slide 21: Calibration

## Slide 23: Windows

## Slide 24: Classification
Far-Wrist Motion
Near-Wrist Motion

## Slide 25: Far Wrist
Synthetic PPG segment based on average of PPG waveform from previously available data

## Slide 28: Near Wrist
Tissue Effect
Venous Blood Movement
Remove using Adaptive and Notch filtering

## Slide 29: Adaptive filtering

## Slide 30: Notch filtering
PPG signal is quasi-periodic
Spectral components around heart rate frequency (HRF) and its harmonics
Suppresses signal around notch frequency and leaves other frequencies as is
Frequency Domain Amplitude Response of a Notch Filter [5]

## Slide 31: Notch filtering - Continued

## Slide 32: Before vs After

## Slide 34: Experimentation and Results

## Slide 35: Setup
8 Subjects:
Rest (reference signal)
Exercise 1 : Walking
Exercise 2 : Slight arm movement
Exercise 3 : Typing
Exercise 4 : Writing

## Slide 36: Results

## Slide 38: Performance Index
  i = Subject Number
  j = Exercise Number
  k = before or after cleansing signal using the described framework
  of kurtosis/skewness spread
  of reference kurtosis/skewness spread

## Slide 39: PI Matrix - Kurtosis

## Slide 40: PI Matrix - Skewness

## Slide 41: Limitations
Accuracy limited with severe motion artifacts over long periods
Synthetic PPG signal relies on quality of previous PPG segment
Requirement of reference signal

## Slide 42: Benefits
Optimized efficiency
Computational complexity reduced
Retains PPG waveform
Live simulation

## Slide 43: Future Work

## Slide 44: Future Recommendations
Collect more data
Overcome limitations of synthetic data
Develop online system
Hardware implementation
Detection of cardiac conditions and health parameters

## Slide 45: Graphical User Interface (GUI)

## Slide 49: Thank You!
Any questions?

## Slide 50: References
[1] S. Naseem, U. K. Khattak, H. Ghazanfar, and A. Irfan, \Prevalence of noncommunicable diseases and their risk factors at a semi-urban community, Pakistan," Pan Afr Med J, vol. 23, p. 151, 2016.
[2] S. Lee, G. Ha, D. Wright, Y. Ma, E. Sen-Gupta, N. Haubrich, P. Branche, W. Li, G. Huppert, M. Johnson, B. Mutlu, K. Li, N. Sheth, J. Wright, y.-s. Huang, M. Mansour, J. Rogers, and R. Ghaari, \Highly exible, wearable, and disposable cardiac biosensors for remote and ambulatory monitoring,“ npj Digital Medicine, vol. 1, 12 2018.
[3] C. Fischer, M. Glos, T. Penzel, and I. Fietze, \Extended algorithm for realtime pulse waveform segmentation and artifact detection in photoplethysmograms,“ Somnologie, vol. 21, pp. 1-11, 05 2017.
[4] Charlton P.H., Villarroel M., Salguiero F. (2016) Waveform Analysis to Estimate Respiratory Rate. In: Secondary Analysis of Electronic Health Records. Springer, Cham. https://doi.org/10.1007/978-3-319-43742-2_26
[5] Wang, Min; Li, Zhe; Zhang, Qirui; Wang, Guoxing. 2019. "Removal of Motion Artifacts in Photoplethysmograph Sensors during Intensive Exercise for Accurate Heart Rate Calculation Based on Frequency Estimation and Notch Filtering" Sensors 19, no. 15: 3312. https://doi.org/10.3390/s19153312